## CCM plasma
Example Template - Project QC Sample Reporting
QC Metrics
| Category | Intended Purpose | Used in this study? | Sub-category | Technical Description | Quantity for this study | Analyzed per batch? | Quantity for each batch | Description of use | Description of Metric | Nature of Metric | Result of Metric Analysis | Acceptance Criteria | Acceptance Results | Criteria Failure Action
| Blanks | To demonstrate process purity during analytical runs, where extraneous signals are identified and accounted for. | No | Solvent blank | MSTFA containing alkane mixture | 52 | Two batches | 26 | Targeted peaks were checked for presence in solvent blanks. | Not analyzed | Quantitative |  | less then 15% | pass | No action required.
| System suitability test (SST) sample(s) | To check that the analytical system is working appropriately and fit-for-purpose. | Yes | Liver samples used for equilibration | Instrument tune followed by analysis of mouse liver 50 mg/mL with 9 alkanes | 11 | Two batches | 4 and 7 | Tune run before analysis. Alkanes in mouse liver were monitored for presence, intensity and stability. | Not analyzed | Quantitative |  | Alkane 32 must lie above an area of 1e+6 with symetrical peak shape Alkane 32 must be stable at least within the last 3 equilibration samples (RSD <15%) | Passed | No action was necessary, but where SST fails, batch is excluded from further analysis.
| Replicate analysis sample(s) | To demonstrate the reproducibility (precision and/or accuracy) of instrumental analysis on the experimental samples or QC samples in the project. | Yes | Intra-study pooled QC samples | All samples that were analyzed were pooled after extraction and the same amount of µL as for samples was used for drying | 11 | Two batches | 5 and 6 | Check for batch drift; check internal standard behaviour. | Checked at data analysis stage. | Quantitative | Alkane 32 has a mean area of 1.72e+8 RSD 9% Cinnamic acid ISTD RSD 12% for batch 1 Alkane 32 has a mean area of 1.58e+8 RSD 10% Cinnamic acid ISTD RSD 11% for batch 2 | Alkane 32 must lie above an area of 1e+6 with symetrical peak shape Alkane 32 must be stable (RSD <15%) Cinnamic acid ISTD must be stable (RSD <15%) | Passed | No action required.
| Internal standards | to demonstrate the technical variability of the analysis INDEPENDENT of the extraction and derivatization of samples | Yes | Alkane mixture Cinnamic acid | All samples had 20 µL/1 mL MSTFA of alkane mixture added during derivatization process. Alkanes are not subject to derivatization. | 18 | Two batches | 9 | Assessed over batch for stability and presence. | Checked at data analysis stage. | Quntitative | Alkane 18 RSD 7% Cinnamic acid RSD 12% for batch 1 Alkane 32 RSD 10% Cinnamic acid RSD 14% for batch 2 | RSD should be <15%, can go up to 20% | Passed | No action required
| Reference material sample(s) | To provide evidence for intra-laboratory or inter-laboratory analytical comparability | No | • Certified Reference Material (CRM) or • Standard Reference Material (SRM™) or • Pseudo Matrix-matched CRM or SRM or • Long-term intra-laboratory sample or • Long-term inter-laboratory sample or • Other (describe) | Text explaining details about the sample, including preparation, sources, concentrations, etc. | (Total number of samples used) | Yes or No | (Number per batch) | Text explaining when this analysis is performed, special conditions required, conditions for exceptions and other operational considerations for this type of sample. | Text explaining details about the data assessment metric. Indicate when the metric is assesed (e.g. during the run or post-analysis assessment). | • Qualitative or • Quantitative | (numerical value or descriptive text) | Text explaining details about the acceptance criteria. | • Passed or • Minimally successful or • Failure Rate | Text explaining actions taken upon unacceptable results.
| Platform specific samples | To provide specific assessment of data quality for specific platforms, possibly recommended by instrument vendors. | No | • Test mixture(s) or • Other (describe) | Text explaining details about the sample, including preparation, sources, concentrations, etc. | (Total number of samples used) | Yes or No | (Number per batch) | Text explaining when this analysis is performed, special conditions required, conditions for exceptions and other operational considerations for this type of sample. | Text explaining details about the data assessment metric. Indicate when the metric is assesed (e.g. during the run or post-analysis assessment). | • Qualitative or • Quantitative | (numerical value or descriptive text) | Text explaining details about the acceptance criteria. | • Passed or • Minimally successful or • Failure Rate | Text explaining actions taken upon unacceptable results.

## CCM urine
Example Template - Project QC Sample Reporting
QC Metrics
| Category | Intended Purpose | Used in this study? | Sub-category | Technical Description | Quantity for this study | Analyzed per batch? | Quantity for each batch | Description of use | Description of Metric | Nature of Metric | Result of Metric Analysis | Acceptance Criteria | Acceptance Results | Criteria Failure Action
| Blanks | To demonstrate process purity during analytical runs, where extraneous signals are identified and accounted for. | No | Solvent blank | MSTFA containing alkane mixture | 46 | Two batches | 23 | Targeted peaks were checked for presence in solvent blanks. | Not analyzed | Quantitative |  | less then 15% | pass | No action required.
| System suitability test (SST) sample(s) | To check that the analytical system is working appropriately and fit-for-purpose. | Yes | Liver samples used for equilibration | Instrument tune followed by analysis of mouse liver 50 mg/mL with 9 alkanes | 11 | Two batches | 4 and 7 | Tune run before analysis. Alkanes in mouse liver were monitored for presence, intensity and stability. | Not analyzed | Quantitative |  | Alkane 32 must lie above an area of 1e+6 with symetrical peak shape Alkane 32 must be stable at least within the last 3 equilibration samples (RSD <15%) | Passed | No action was necessary, but where SST fails, batch is excluded from further analysis.
| Replicate analysis sample(s) | To demonstrate the reproducibility (precision and/or accuracy) of instrumental analysis on the experimental samples or QC samples in the project. | Yes | Intra-study pooled QC samples | All samples that were analyzed were pooled after extraction and the same amount of µL as for samples was used for drying | 19 | Two batches | 10 and 9 | Check for batch drift; check internal standard behaviour. | Checked at data analysis stage. | Quantitative | Alkane 32 has a mean area of 1.61e+8 RSD 37% Cinnamic acid ISTD RSD 27% for batch 1 Alkane 32 has a mean area of 1.74e+8 RSD 32% Cinnamic acid ISTD RSD 21% for batch 2 | Alkane 32 must lie above an area of 1e+6 with symetrical peak shape Alkane 32 must be stable (RSD <15%) Cinnamic acid ISTD must be stable (RSD <15%) | drift observed during run | results should be carefully when interpretate
| Internal standards | to demonstrate the technical variability of the analysis INDEPENDENT of the extraction and derivatization of samples | Yes | Alkane mixture Cinnamic acid | All samples had 20 µL/1 mL MSTFA of alkane mixture added during derivatization process. Alkanes are not subject to derivatization. | 18 | Two batches | 9 | Assessed over batch for stability and presence. | Checked at data analysis stage. | Quntitative | Alkane 32 RSD 35% Cinnamic acid RSD 32% for batch 1 Alkane 22 RSD 21% Cinnamic acid RSD 26% for batch 2 | RSD should be <15%, can go up to 20% | drift observed during run | results should be carefully interpretated
| Reference material sample(s) | To provide evidence for intra-laboratory or inter-laboratory analytical comparability | No | • Certified Reference Material (CRM) or • Standard Reference Material (SRM™) or • Pseudo Matrix-matched CRM or SRM or • Long-term intra-laboratory sample or • Long-term inter-laboratory sample or • Other (describe) | Text explaining details about the sample, including preparation, sources, concentrations, etc. | (Total number of samples used) | Yes or No | (Number per batch) | Text explaining when this analysis is performed, special conditions required, conditions for exceptions and other operational considerations for this type of sample. | Text explaining details about the data assessment metric. Indicate when the metric is assesed (e.g. during the run or post-analysis assessment). | • Qualitative or • Quantitative | (numerical value or descriptive text) | Text explaining details about the acceptance criteria. | • Passed or • Minimally successful or • Failure Rate | Text explaining actions taken upon unacceptable results.
| Platform specific samples | To provide specific assessment of data quality for specific platforms, possibly recommended by instrument vendors. | No | • Test mixture(s) or • Other (describe) | Text explaining details about the sample, including preparation, sources, concentrations, etc. | (Total number of samples used) | Yes or No | (Number per batch) | Text explaining when this analysis is performed, special conditions required, conditions for exceptions and other operational considerations for this type of sample. | Text explaining details about the data assessment metric. Indicate when the metric is assesed (e.g. during the run or post-analysis assessment). | • Qualitative or • Quantitative | (numerical value or descriptive text) | Text explaining details about the acceptance criteria. | • Passed or • Minimally successful or • Failure Rate | Text explaining actions taken upon unacceptable results.

## tryptophan
Example Template - Project QC Sample Reporting
QC Metrics
| Category | Intended Purpose | Used in this study? | Sub-category | Technical Description | Quantity for this study | Analyzed per batch? | Quantity for each batch | Description of use | Description of Metric | Nature of Metric | Result of Metric Analysis | Acceptance Criteria | Acceptance Results | Criteria Failure Action
| Blanks | To demonstrate process purity during analytical runs, where extraneous signals are identified and accounted for. | Yes | • True Blank used | solvent blanks were used for wash out of carry over peaks of highly concentrated calibrants and | 80 solvent blanks | Yes, analysed per batch in 2 batches | 40 | blanks were used to assess the LOD and LLOQ and to wash off carry over peaks that occured after the highest calibration standard for some metabolites | blanks were used to assess LOD and LLOQ | Quantitative | numerical value | LOD threshhold is 3x blank median | Passed | Column check and flush and run new solvents to remove the background noise
| System suitability test (SST) sample(s) | To check that the analytical system is working appropriately and fit-for-purpose. | No | Calibration sample | 5 equillibration samples were run in the beginning of a sequence to test the system | 5 | Yes | 5 | These equillibration samples were just evlauated in peak shape and intensity to test if the system is stable | These equillibration samples were just evlauated in peak shape and intensity to test if the system is stable | • Qualitative | value of the RSD fine for most metabolites | RSD of the last 3 replicates of <15% | Passed | Running more replicates of the equillibration samples to stabilize the system
| Replicate analysis sample(s) | To demonstrate the reproducibility (precision and/or accuracy) of instrumental analysis on the experimental samples or QC samples in the project. | Yes | • Pooled Sample and • Reference Standard | after the extraction 20µl of all samples were pooled to one QC and injected along the sequence | 16 pooled QC samples, 18 referecne QC samples | Yes | 8 pooled QC and 9 reference QC | assessment of intra-batch and inter-batch variation of replictaes alonge the sequence, every loop of 6 samples was followed by QC samples throughout the sequence |  | • Qualitative or • Quantitative | intra-batch RSD fine for all metabolites, inter-batch RSD fine for most metabolites. | RSD should be <15%, can go up to 20% | Passed for most metabolites only | Batch correction done
| Internal standards (IS) | to demonstrate the technical variability of the analysis indepndent of the extraction and of samples | Yes | 13C or deuterated metabolites:melatonin, tryptophan, hydroxykynurenine, kynurenine, idoxylsulfate, hydroxyindoleacetate, | all internal standards were added to the extraction solvent in a final concentration of 0.1µg/ml |  | Two batches | equal concentration in all samples, QCs and calibrants | Assessed over batch for stability and presence of matrix effects and ion suppression. | Checked at data analysis stage. | Quntitative | all IS showed matrix effect and some degree of ion suppression | RSD should be <15%, can go up to 20% | passed, no drift observed during run | normalization to IS
| Reference material sample(s) | To provide evidence for intra-laboratory or inter-laboratory analytical comparability | No | • Certified Reference Material (CRM) or • Standard Reference Material (SRM™) or • Pseudo Matrix-matched CRM or SRM or • Long-term intra-laboratory sample or • Long-term inter-laboratory sample or • Other (describe) | Text explaining details about the sample, including preparation, sources, concentrations, etc. | (Total number of samples used) | No | (Number per batch) | Text explaining when this analysis is performed, special conditions required, conditions for exceptions and other operational considerations for this type of sample. | Text explaining details about the data assessment metric. Indicate when the metric is assesed (e.g. during the run or post-analysis assessment). | • Qualitative or • Quantitative | (numerical value or descriptive text) | Text explaining details about the acceptance criteria. | • Passed or • Minimally successful or • Failure Rate | Text explaining actions taken upon unacceptable results.
| Platform specific samples | To provide specific assessment of data quality for specific platforms, possibly recommended by instrument vendors. | No | • Test mixture(s) or • Other (describe) | Text explaining details about the sample, including preparation, sources, concentrations, etc. | (Total number of samples used) | No | (Number per batch) | Text explaining when this analysis is performed, special conditions required, conditions for exceptions and other operational considerations for this type of sample. | Text explaining details about the data assessment metric. Indicate when the metric is assesed (e.g. during the run or post-analysis assessment). | • Qualitative or • Quantitative | (numerical value or descriptive text) | Text explaining details about the acceptance criteria. | • Passed or • Minimally successful or • Failure Rate | Text explaining actions taken upon unacceptable results.

## Biocrates
Example Template - Project QC Sample Reporting
QC Metrics
| Category | Intended Purpose | Used in this study? | Sub-category | Technical Description | Quantity for this study | Analyzed per batch? | Quantity for each batch | Description of use | Description of Metric | Nature of Metric | Result of Metric Analysis | Acceptance Criteria | Acceptance Results | Criteria Failure Action
| Blanks | To demonstrate process purity during analytical runs, where extraneous signals are identified and accounted for. | Yes | • True Blank used | solvent blanks were used for evaluation of the background | 10 solvent blanks | Yes, analysed per batch in 2 batches | 5 | blanks were used to assess the LOD and LLOQ and to wash off carry over peaks that occured after the highest calibration standard for some metabolites | blanks were used to assess LOD and LLOQ | Quantitative | numerical value | LOD threshhold is 3x blank median | Passed | Column check and flush and run new solvents to remove the background noise
| System suitability test (SST) sample(s) | To check that the analytical system is working appropriately and fit-for-purpose. | Yes | Calibration test sample for LC and FIA method separately | LC and FIA method were run in positive and negative mode with the SST test sample to check and correct retention times of the LC method and stable flow and intensity for the method | 6 | Yes | 3 | LC and FIA method run in positive and negative mode with the SST test sample to check and correct retention times of the LC methods | These SST samples were used for evlauation of peak shape and intensity to test if the system and retention times are stable | Qualitative | value of the RSD fine for all metabolites | RSD of the 3 replicates of <15% | Passed | Running more replicates of the equillibration samples to stabilize the system
| Replicate analysis sample(s) | To demonstrate the reproducibility (precision and/or accuracy) of instrumental analysis on the experimental samples or QC samples in the project. | Yes | • Pooled Sample and • Reference Standard | after the extraction 20µl of all samples were pooled to one QC and injected along the sequence | 10 pooled QC samples, 10 reference QC samples | Yes | 5 pooled QC and 5 reference QC | assessment of intra-batch and inter-batch variation of replictaes alonge the sequence, every loop of 6 samples was followed by QC samples throughout the sequence | method is semi-quantitative per 96 well plate | Quantitative | intra-batch RSD and inter-batch RSD fine for all metabolites. | RSD should be <15%, can go up to 20% | Passed for | no action needed
| Internal standards (IS) | to demonstrate the technical variability of the analysis independent of the extraction and of samples | Yes | 13C or deuterated metabolites | all internal standards were added to the 96 well plate prior to extraction | 96 | Two batches | equal concentration in all samples, QCs and calibrants | Assessed in the software and used for IS normalization prior to calibration | Checked at data analysis stage. | Quntitative | no batch effect observed | RSD should be <15%, can go up to 20% | passed, no drift observed during run | no action needed
| Reference material sample(s) | To provide evidence for intra-laboratory or inter-laboratory analytical comparability | No | • Certified Reference Material (CRM) or • Standard Reference Material (SRM™) or • Pseudo Matrix-matched CRM or SRM or • Long-term intra-laboratory sample or • Long-term inter-laboratory sample or • Other (describe) | Text explaining details about the sample, including preparation, sources, concentrations, etc. | (Total number of samples used) | No | (Number per batch) | Text explaining when this analysis is performed, special conditions required, conditions for exceptions and other operational considerations for this type of sample. | Text explaining details about the data assessment metric. Indicate when the metric is assesed (e.g. during the run or post-analysis assessment). | • Qualitative or • Quantitative | (numerical value or descriptive text) | Text explaining details about the acceptance criteria. | • Passed or • Minimally successful or • Failure Rate | Text explaining actions taken upon unacceptable results.
| Platform specific samples | To provide specific assessment of data quality for specific platforms, possibly recommended by instrument vendors. | No | • Test mixture(s) or • Other (describe) | Text explaining details about the sample, including preparation, sources, concentrations, etc. | (Total number of samples used) | No | (Number per batch) | Text explaining when this analysis is performed, special conditions required, conditions for exceptions and other operational considerations for this type of sample. | Text explaining details about the data assessment metric. Indicate when the metric is assesed (e.g. during the run or post-analysis assessment). | • Qualitative or • Quantitative | (numerical value or descriptive text) | Text explaining details about the acceptance criteria. | • Passed or • Minimally successful or • Failure Rate | Text explaining actions taken upon unacceptable results.